## Updated products
Medical Device | Technical description | Norms and Standards
Face Shield | Made of clear plastic and providing good visibility to both the wearer and the patient Adjustable band to attach firmly around the head and fit snuggly against the forehead, fog-resistant (preferable). Completely covers the sides and length of the face. May be re-usable (made of robust material which can be cleaned and disinfected) or disposable. | • EU PPE Regulation 2016/425
• EN 166
• ANSI/ISEA Z87.1 or equivalent
Gown | Single-use, length mid-calf. | • EU PPE Regulation 2016/425 and EU MDD Directive 93/42/EEC
• FDA Class I or II medical device, or equivalent
• EN 13795 any performance level, or
• AAMI PB70 all levels acceptable, or equivalent
Goggles, protective | Good seal with the skin of the face, flexible PVC frame to easily fit all face contours with even pressure, enclose eyes and the surrounding areas, accommodate wearers with prescription glasses; clear plastic lens with fog- and scratch-resistant treatments; adjustable band to secure firmly so as not to become loose during clinical activity; indirect venting to avoid fogging. May be re-usable (provided appropriate arrangements for decontamination are in place) or disposable. | • EU PPE Regulation 2016/425
• EN 166
• ANSI/ISEA Z87.1 or equivalent
Gloves, examination, non-sterile | Gloves, examination, nitrile, powder-free, non-sterile, single-use. | • EU MDD Directive 93/42/EEC Category III
| Gloves should have long cuffs, reaching well above the wrist, ideally to mid-forearm. | • EU PPE Regulation 2016/425 Category III
| Sizes: small, medium, large. | • EN 455
• EN 374
• ANSI/ISEA 105
• ASTM D6319 or equivalent
Mask, surgical – healthcare worker | Surgical mask, good breathability; internal and external faces should be clearly identified | • EU MDD Directive 93/42/EEC Category III or equivalent
| Type II or higher | • EN 14683 Type II, IR, IIIR
• ASTM F2100 minimum level 1 or equivalent
Mask, surgical – patient | Surgical mask, good breathability; internal and external faces should be clearly identified | • EN 14683 any type including Type I
| Type I. | • ASTM F2100 minimum level 1 or equivalent
Particulate respirator, grade N95 or higher | N95 or FFP2 respirator, or higher | • Minimum "N95" respirator according to FDA Class II, under 21 CFR 878.4040, and CDC NIOSH, or
| Good breathability with a design that does not collapse against the mouth (e.g. duckbill, cupshaped). | • Minimum "FFP2 according to EN 149, EU PPE
• Regulation 2016/425 Category III, or equivalent
Source: These technical requirements were extracted from the Disease Commodity Package COVID-19 v4 developed by the World Health Organization (WHO).

## Tecnical Propousal N95
1: Proof of eligibility | A formal statement must be submitted by the supplier confirming that the N95 Respirator offered is in compliance with the following technical requirements: • The product must be in alignment with the technical description and at least one of the norms or standards (indicating which one) • The product must be in conformity with the regulatory requirements in the country of origin and at least one of the members of the International Medical Device Regulators Forum (IMDRF); and • The product must be manufactured in compliance with ISO 13485: Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes, or an equivalent quality system. • The supplier must present proof of compliance of the product with at least one of the correspondent technical standards. (certificates must be issued by the certification bodies or the test reports in accordance with the International Laboratory Accreditation Cooperation (ILAC) laboratory standards, or equivalent)
2: Product description | The product description must be provided by the supplier. It must include a) brand or trade name, b) product descriptive information and intended use, c) a list of the materials, d) expected lifetime, and e) approved storage conditions..
3: Manufacturing information | 1. List of all sites involved in the manufacturing process, including name, current address and manufacturing responsibilities/activities for each site as well as the list of outsourced processes with direct product impact, including details of the supplier for each process. 2. Manufacturing site license issued according to the legislation of the country of origin. 3. Proof of compliance of each manufacturer site involved in the manufacturing process with ISO 13485: Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes, an equivalent quality system standard; or a certificate of Good Manufacturing Practice (GMP), issued by the National Regulatory Authority of the country of manufacturer; or a Medical Device Single Audit Program (MDSAP) certificate.
4: Product labelling | The supplier must provide a copy of the product labelling. This includes the label, instructions for use, and information related to the identification, technical description, intended purpose and proper use of the product, and other instructional materials provided for users of the product, both professional and lay persons, as appropriate, and for relevant third parties*. - Labels must minimally include a) brand or trade name of the medical device, b) catalog number, c) contact information of the manufacturer, d) recommendations that would reduce sufficiently the risk of use, e) intended use, f) handling measures for storage and transport, g) warnings and precautions, and h) particular product conditions (e.g. sterility). - The instructions for use, where possible, should comply with the principles of labelling for medical devices and IVD medical devices of IMDRF/GRRP WG/N52 FINAL:2019. * In any of the official languages of the PAHO Region: English, Spanish, French and/or Portuguese.
Right to share information with PAHO Member States PAHO, at its sole discretion, may share with the corresponding National Regulatory Authorities from the recipient country, documents from the technical proposal or other documents and data requested for clarification or to resolve any potential claims or inquiries.

## Tecnical Propousal other produ
1: Proof of eligibility | A formal statement must be submitted by the supplier confirming that each product offered is in compliance with the following technical requirements: • The products must be in alignment with the technical description and at least one of the norms or standards (indicating which one) • The products must be in conformity with the regulatory requirements in the country of origin and at least one of the members of the International Medical Device Regulators Forum (IMDRF); and • The products must be manufactured in compliance with ISO 13485: Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes, or an equivalent quality system
2: Product description | The product description must be provided by the supplier. It must include a) brand or trade name, b) product descriptive information and intended use, c) a list of the materials, d) expected lifetime, and e) approved storage conditions..
3: Manufacturing information | 1. List of all sites involved in the manufacturing process, including name, current address and manufacturing responsibilities/activities for each site as well as the list of outsourced processes with direct product impact, including details of the supplier for each process. 2. Manufacturing site license issued according to the legislation of the country of origin. 3. Proof of compliance of each manufacturer site involved in the manufacturing process with ISO 13485: Medical Devices – Quality Management Systems – Requirements for Regulatory Purposes, an equivalent quality system standard; or a certificate of Good Manufacturing Practice (GMP), issued by the National Regulatory Authority of the country of manufacturer; or a Medical Device Single Audit Program (MDSAP) certificate.
4: Product labelling | The supplier must provide a copy of the product labelling. This includes the label, instructions for use, and information related to the identification, technical description, intended purpose and proper use of the PPE, and other instructional materials provided for users of the PPE, both professional and lay persons, as appropriate, and for relevant third parties*. - Labels must minimally include a) brand or trade name of the medical device, b) catalog number, c) contact information of the manufacturer, d) recommendations that would reduce sufficiently the risk of use, e) intended use, f) handling measures for storage and transport, g) warnings and precautions, and h) particular product conditions (e.g. sterility). - The instructions for use, where possible, should comply with the principles of labelling for medical devices and IVD medical devices of IMDRF/GRRP WG/N52 FINAL:2019. * In any of the official languages of the PAHO Region: English, Spanish, French and/or Portuguese.
Right to share information with PAHO Member States PAHO, at its sole discretion, may share with the corresponding National Regulatory Authorities from the recipient country, documents from the technical proposal or other documents and data requested for clarification or to resolve any potential claims or inquiries.